## CT2BE & CT5BN
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised range 4.2.25 | QC 1 | QC 2 | ? Analyser precision not good enough to achieve the current 5% target SD |  | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 20 | 195.1 | Mean | 19 |  |  | Mean |  | #DIV/0!
|  | Lot No. |  | Month | July | establish current target and range using 5% target SD and complete |  |  | SD | 1 | 9.755 | SD | 0.95 | 0 |  | SD | 0 | #DIV/0!
Level 1 |  | CT2BE |  | Year | 2025 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 18 | 175.59 | lower limit | 17.1 | 0 |  | lower limit | 0 | #DIV/0!
Level 2 |  | CT5BN |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 22 | 214.61 | upper limit | 20.9 | 0 |  | upper limit | 0 | #DIV/0!
Date | QC Level 1 | QC Level 2 | Comments
| 2025-07-10 | 2025-07-03 | [15 empty] | Level 1 | CT2BE |  |  |  | Level 2 | CT5BN
target | 17-21 | 116-271 | [15 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2025-07-03 |  | 177 | Manufacturers range use 116-271 | [14 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
08/07/2025 (am) |  | 202 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
08/07/2025 (pm) |  | 201 |  |  |  | TREND REVIEW | [11 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-10 | 21 | 189 | Manufacturers range 0-40 (use 17-21) | [14 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-15 | 19 | 188 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-17 | 19 | 193 |  |  |  | Range | QC 1 | QC 2 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-22 | 19 | 198 |  |  |  | Mean | 19 | 195.1 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-24 | 18 | 196 |  |  |  | SD | 1 | 8.7743 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-29 | 18 | 199 |  |  |  | % CV | 5.2632 | 4.4974 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
2025-07-31 | 19 | 208 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61
[18 empty] | 19 | 18 | 20 | 17 | 21 | 195.1 | 185.345 | 204.855 | 175.59 | 214.61

## CT2BD & CT5BN
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised range 4.2.25 | QC 1 | QC 2 | ? Analyser precision not good enough to achieve the current 5% target SD |  | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 19.5 | #DIV/0! | Mean | 19.2857 |  |  | Mean |  | #DIV/0!
|  | Lot No. |  | Month | Feb - July | establish current target and range using 5% target SD and complete |  |  | SD | 0.975 | #DIV/0! | SD | 0.96429 | 0 |  | SD | 0 | #DIV/0!
Level 1 |  | CT2BD |  | Year | 2025 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 17.55 | #DIV/0! | lower limit | 17.3571 | 0 |  | lower limit | 0 | #DIV/0!
Level 2 |  | CT5BN |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 21.45 | #DIV/0! | upper limit | 21.2143 | 0 |  | upper limit | 0 | #DIV/0!
Date | QC Level 1 | QC Level 2 | Comments
| 2025-02-25 | 2025-07-01 | [15 empty] | Level 1 | CT2BD |  |  |  | Level 2 | CT5BN
target | 17-21 | [16 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2025-02-25 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-02-27 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-04 | 21 |  |  |  |  | TREND REVIEW | [11 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-06 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-11 | 17 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-11 | [5 empty] | Range | QC 1 | QC 2 | [9 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | Mean | 19.2683 | 193.3333 | [9 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-13 | 18 |  |  |  |  | SD | 1.5973 | 14.1539 | [9 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-18 | 18 |  |  |  |  | % CV | 8.2895 | 7.321 | [9 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-20 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-25 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-03-27 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-01 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-03 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-08 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-10 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-15 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-17 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-22 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-24 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-04-29 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-01 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-06 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-08 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-13 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-15 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-20 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-22 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-27 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-05-29 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-03 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-05 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-10 | 17 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-12 | 17 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-17 | 17 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-19 | 16 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-19 | 23 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-26 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-06-26 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-07-01 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2025-07-03 | 17 | 177 | NEW LOT QC IN USE - To assign ranges | [14 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
08/07/2025 (am) | 20 | 202 | [15 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
08/07/2025 (pm) | 22 | 201 | QC 1 accepted by RH | [14 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[18 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[18 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[18 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[18 empty] | 19 | 18 | 20 | 17 | 21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## CT2BD & CT5BM
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised range 4.2.25 | QC 1 | QC 2 | ? Analyser precision not good enough to achieve the current 5% target SD |  | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 19.5 | 206.5556 | Mean | 19.2857 |  |  | Mean |  | #DIV/0!
|  | Lot No. |  | Month | April | establish current target and range using 5% target SD and complete |  |  | SD | 0.975 | 10.3278 | SD | 0.96429 | 0 |  | SD | 0 | #DIV/0!
Level 1 |  | CT2BD |  | Year | 2025 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 17.55 | 185.9 | lower limit | 17.3571 | 0 |  | lower limit | 0 | #DIV/0!
Level 2 |  | CT5BM |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 21.45 | 227.2111 | upper limit | 21.2143 | 0 |  | upper limit | 0 | #DIV/0!
Date | QC Level 1 | QC Level 2 | Comments
| 2025-02-25 | 2025-04-22 | [15 empty] | Level 1 | CT2BD |  |  |  | Level 2 | CT5BM
target | 17-21 | 186-227 | [15 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2025-02-25 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-02-27 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-04 | 21 |  |  |  |  | TREND REVIEW | [11 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-06 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-11 | 17 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-11 | [5 empty] | Range | QC 1 | QC 2 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[6 empty] | Mean | 19.1795 | 200.087 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-13 | 18 |  |  |  |  | SD | 1.5706 | 13.3618 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-18 | 18 |  |  |  |  | % CV | 8.1892 | 6.678 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-20 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-25 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-03-27 | 20 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-01 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-03 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-08 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-10 | 18 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-15 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-17 | 18 |  | 186-227 | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-22 | 18 | 204 | New LOT Pos QC - Same ranges | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-24 | 18 | 205 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-04-29 | 18 | 201 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-01 | 19 | 196 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-06 | 20 | 204 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-08 | 21 | 211 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-13 | 20 | 200 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-15 | 20 | 211 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-20 | 21 | 227 | running on manufacturers ranges | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-22 | 20 | 213 | New QC range assigned 186-227 | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-27 | 21 | 197 | calibration due 29.05.25 | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-05-29 | 20 | 189 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-03 | 18 | 198 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-05 | 20 | 215 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-10 | 17 | 217 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-12 | 17 | 180 | repeat 179 - to recalibrate | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-17 | 17 | 197 | calibrated | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-19 | 16 | 178 | RH/HG Extracted IQC okay. Small variation. Run accepted | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-19 | 23 | 182 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-26 | 20 | 188 | calibrated | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-06-26 | 18 | 193 | calibrated | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-07-01 | 19 | 216 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
2025-07-03 | 17 | 180 | new LOT in use (CT5BN). High QC accepted as new lot and extracted IQC all acceptable | [14 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[18 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[18 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[18 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[18 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[18 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111
[18 empty] | 19 | 18 | 20 | 17 | 21 | 206.5556 | 196.2278 | 216.8833 | 185.9 | 227.2111

## CT2BD & CT5BL
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised range 4.2.25 | QC 1 | QC 2 | ? Analyser precision not good enough to achieve the current 5% target SD |  | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 19.5 | 202.375 | Mean | 19.2857 |  |  | Mean |  | #DIV/0!
|  | Lot No. |  | Month | February | establish current target and range using 5% target SD and complete |  |  | SD | 0.975 | 10.1188 | SD | 0.96429 | 0 |  | SD | 0 | #DIV/0!
Level 1 |  | CT2BD |  | Year | 2025 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 17.55 | 182.1375 | lower limit | 17.3571 | 0 |  | lower limit | 0 | #DIV/0!
Level 2 |  | CT5BL |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 21.45 | 222.6125 | upper limit | 21.2143 | 0 |  | upper limit | 0 | #DIV/0!
Date | QC Level 1 | QC Level 2 | Comments
| 2025-02-25 | 2025-04-03 | [15 empty] | Level 1 | CT2BD |  |  |  | Level 2 | CT5BL
target | 17-21 | 182-223 | L | [14 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2025-02-25 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-02-27 | 21 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-04 | 21 |  |  |  |  | TREND REVIEW | [11 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-06 | 19 | [16 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-11 | 17 |  | calibrated | [14 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-11 |  | 201 | New LOT CT5BLtarget 194 manufacturers range: 116-271 |  |  | Range | QC 1 | QC 2 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
|  |  | Range assigned for CT2BD 17-21 |  |  | Mean | 19.375 | 201.4167 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-13 | 18 | 204 |  |  |  | SD | 1.3601 | 6.3024 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-18 | 18 | 188 |  |  |  | % CV | 7.0201 | 3.129 | [9 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-20 | 19 | 202 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-25 | 20 | 207 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-03-27 | 20 | 201 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-04-01 | 21 | 207 | Borderline high QC, not clinically significant, Run accepted HG | [14 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-04-03 | 21 | 209 | New range assigned to Lot CT5BL 182-223 | [14 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-04-08 | 19 | 191 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-04-10 | 18 | 199 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-04-15 | 19 | 205 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
2025-04-17 | 18 | 203 | [15 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125
[18 empty] | 19 | 18 | 20 | 17 | 21 | 202.375 | 192.2562 | 212.4938 | 182.1375 | 222.6125

## LOT CT2BC(1) and CT5BK(2)
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised range 4.2.25 | QC 1 | QC 2 | ? Analyser precision not good enough to achieve the current 5% target SD |  | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 20.8333 | 185.5 | Mean | 20.8462 |  |  | Mean |  | #DIV/0!
|  | Lot No. |  | Month | DECEMBER/JANUARY | establish current target and range using 5% target SD and complete |  |  | SD | 1.0417 | 9.275 | SD | 1.0423 | 0 |  | SD | 0 | #DIV/0!
Level 1 |  | CT2BC |  | Year | 2024/2025 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 18.75 | 166.95 | lower limit | 18.7615 | 0 |  | lower limit | 0 | #DIV/0!
Level 2 |  | CT5BK |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 22.9167 | 204.05 | upper limit | 22.9308 | 0 |  | upper limit | 0 | #DIV/0!
Date | QC Level 1 | QC Level 2 | Comments
| 2025-02-04 | 2024-12-24 | [15 empty] | Level 1 | CT2BC |  |  |  | Level 2 | CT5BK
target | 19-23 | 167-204 | [15 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2024-12-12 | 23 | 182 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2024-12-17 | 23 | 182 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2024-12-19 | 22 | 173 |  |  |  | TREND REVIEW | [11 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2024-12-24 | 22 | 185 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
|  |  | Targets to be assigned and given to immunology for lvl 2 | [14 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-02 | 22 | 182 |  |  |  | Range | QC 1 | QC 2 | [9 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-07 | 18 | 196 | Lvl1 accepted RH (calibrated today) |  |  | Mean | 20.5263 | 186.4783 | [9 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-09 | 18 | 195 | Lvl 1 accepted CT (to calibrate again next time) |  |  | SD | 1.5044 | 8.8875 | [9 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-14 | 21 | 195 | QC1 accepted by CT (Calibrated 14/01/25) |  |  | % CV | 7.329 | 4.766 | [9 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-16 | 20 | 198 | QC1 accepted by RH | [14 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-21 | 21 | 192 | QC 1 low - Discussed with SH and accepted | [14 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-23 | 19 | 191 | QC 1 low - Discussed with SH and accepted |  |  | Range | QC 1 | QC 2 | [9 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-25 | 21 | 190 | QC 1 LOW - Accepted by CT |  |  | Mean | [11 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-01-28 | 21 | 190 |  |  |  | SD | [11 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
|  |  | new target range for QC1 19-23 |  |  | %CV | [11 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-04 | 20 | 189 | [15 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-06 | 20 | 199 | [15 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-11 | 19 | 171 | calibrated 11/02/25 | [14 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-13 | 20 | 181 | [15 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-18 | 19 | 188 | [15 empty] | 21 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-20 | 21 | 200 | [15 empty] | 24 | 20 | 22 | 19 | 23 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-25 |  | 175 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-02-27 |  | 177 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-03-04 |  | 186 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
2025-03-06 |  | 172 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 176.225 | 194.775 | 166.95 | 204.05
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 0 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.5 | 0 | 0 | 0 | 0

## LOT CT2BC and CT5BJ
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised Range 31/10/24 | QC 1 | QC 2 | ? Analyser precision not good enough to achieve the current 5% target SD | Revised Range 12/11/24 | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 20.2 | 192.8571 | Mean | 22 | 185.2308 |  | Mean | 24 | #DIV/0!
|  | Lot No. |  | Month | OCTOBER/NOVEMBER | establish current target and range using 5% target SD and complete |  |  | SD | 1.01 | 9.6429 | SD | 1.1 | 9.2615 |  | SD | 1.2 | #DIV/0!
Level 1 |  | CT2BC |  | Year | 2024 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 18.18 | 173.5714 | lower limit | 19.8 | 166.7077 |  | lower limit | 21.6 | #DIV/0!
Level 2 |  | CT2BJ |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 22.22 | 212.1429 | upper limit | 24.2 | 203.7538 |  | upper limit | 26.4 | #DIV/0!
Date | QC Level 1 |  | Comments
| 2024-11-12 | 2024-10-31 | Immunology ranges | [14 empty] | Level 1 | CT2BC |  |  |  | Level 2 | CT2BJ
target | 22-26 | 167-204 | [15 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
[18 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-01 |  | 194 | [15 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-03 |  | 188 |  |  |  | TREND REVIEW | [11 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-08 | 20 | 205 | [15 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-10 | 20 | 193 | [15 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-15 | 23 | 193 | QC1 accepted by CT |  |  | Range | QC 1 | QC 2 | [9 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-17 | 17 | 183 | Still running on temporary range 17-21, accepted |  |  | Mean | 22.7826 | 185.4762 | [9 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-22 | 21 | 194 | New ranges assigned |  |  | SD | 2.1096 | 8.3763 | [9 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-24 | 20 | 185 |  |  |  | % CV | 9.2599 | 4.5161 | [9 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-29 | 23 | 177 | calibrated 29/10/2024 |  | IQC accepted as only borderline high and extracted IQC within targets. Calibrate again before next run. HG | [12 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-29 | 25 |  | Repeat as low IQC high | [14 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-10-31 | 24 | 185 | Acceptd by biochemist Richard. New QC lot mean revised |  |  | Range | QC 1 | QC 2 | [9 empty] | 22 | 21 | 23 | 20 | 24 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-11-05 | 22 | 180 |  |  |  | Mean | [11 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-11-05 | 25 |  |  |  |  | SD | [11 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-11-07 | 26 | 191 | Accepted 06/11/24 as extracted IQC okay. 05/11/24 to re-extract. HG |  |  | %CV | [11 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 192.8571 | 183.2143 | 202.5 | 173.5714 | 212.1429
2024-11-12 | 23 | 173 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-11-12 |  | 172 | New ranges assigned for QC 2 | [14 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-11-14 | 24 | 192 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-11-19 | 25 | 175 | Lvl 1 QC accepted RJH | [14 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-11-27 | 25 | 180 | Lvl 1 QC accepted HG | [14 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-11-28 | 23 | 183 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-12-03 | 22 | 190 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-12-05 | 24 | 185 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-12-11 | 24 | 177 | [15 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-12-12 | 23 | [16 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-12-17 | 23 | [16 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
2024-12-19 | 22 | [16 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 24 | 22.8 | 25.2 | 21.6 | 26.4 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 22 | 21 | 23 | 20 | 24 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 185.2308 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538
[18 empty] | 19 | 18 | 20 | 17 | 21 | 192.8571 | 175.9692 | 194.4923 | 166.7077 | 203.7538

## LOT CT2BB and CT5BH
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Revised Range | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 20.4 | 197.2 | Mean | 19
|  | Lot No. |  | Month | APRIL-OCTOBER | establish current target and range using 5% target SD and complete |  |  | SD | 1.02 | 9.86 | SD | 0.95
Level 1 |  | CT2BB |  | Year | 2024 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 18.36 | 177.48 | lower limit | 17.1
Level 2 |  | CT5BH |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 22.44 | 216.92 | upper limit | 20.9
Phadia QC (end of run)
Date | QC Level 1 | QC Level 2 | Comments
| 17-21 | 177-217 | Immunology ranges | [8 empty] | Date | QC Level 1 | QC Level 2 | Level 1 | CT2BB |  |  |  | Level 2 | CT5BH
target | 19 | 197 | [12 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2024-04-18 | 22 | 191 | [9 empty] | 2024-04-16 | 19 | 213 | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-04-23 | 22 | 195 | [9 empty] | 2024-04-18 | 23 | 199 | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-04-25 | 18 | 224 |  |  | shift upward QC2, ask for calibration | TREND REVIEW | [6 empty] | 22 | 205 | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-04-30 | 18 | 187 | [9 empty] | 2024-04-30 | 20 | 209 | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-02 | 22 | 189 |  |  | fcal | [6 empty] | 2024-05-07 | 22 | 195 | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
07.05.24 | 20 | 198 |  |  |  | Range | QC 1 | QC 2 |  |  |  | 2024-05-14 | 18 | 190 | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
07.05.24 | 22 | 195 |  |  |  | Mean | 19.5102 | 195.617 | [6 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-09 | 24 | 210 |  |  | QC accepted, read note. | SD | 1.6218 | 10.1757 | [6 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-14 | 22 | 194 |  |  |  | % CV | 8.3124 | 5.2019 | [6 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-16 | 20 | 178 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-21 | 19 | 199 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-23 | 20 | 179 |  |  |  | Range | QC 1 | QC 2 | [6 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-28 | 20 | 189 |  |  |  | Mean | [8 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-30 | 23 | 190 |  |  | QC accepted as extracted QC ok. Will be calibrated before next run. | SD | [8 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-04 | 18 | 198 |  |  |  | %CV | [8 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-06 | 21 | 198 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-11 | 18 | 186 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-13 | 19 | 187 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-18 | 18 | 183 |  |  | Lvl 2 rerun | [9 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-20 | 17 | 187 |  |  | Lvl 2 rerun, Lvl 1 QC accepted. Cal'd before next run. | [9 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-05-25 | 19 | 201 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-06-27 | 19 | 204 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-02 | 18 | 181 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-04 | 20 | 178 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-09 | 19 | 189 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-11 | 18 | 201 | [12 empty] | 20 | 19 | 21 | 18 | 22 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-16 | 18 | 197 |  |  | Revised target for L1 based on results during weekly calibration for 3 weeks | [9 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-18 | 21 | 187 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-23 | 21 | 201 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-25 | 20 | 190 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-07-30 | 18 | 189 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-01 | 18 | 185 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-06 | 21 | 192 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-08 | 17 | 195 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-13 | 18 | 211 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-15 | 19 | 208 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-20 | 19 | 213 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-22 | 19 | 200 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-27 | 19 | 211 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-08-29 | 19 | 203 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-04 | 19 | 199 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-05 | 19 | 190 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-10 | 18 | 204 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-12 | 21 | 199 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-19 | 20 | 211 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-24 | 19 | 191 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-09-26 | 21 | 207 | [12 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
2024-10-01 | 18
2024-10-03 | 18 | [13 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
[15 empty] | 19 | 18 | 20 | 17 | 21 | 197.2 | 187.34 | 207.06 | 177.48 | 216.92
Note on F18: Hall Sebastian: Level 1 QC not rerun by immunology as they were still using the manufacturers QC ranges due to biochem not communicating the new ranges. This has now been communicated. QC accepted as only 1 level, still within 3SD and extracted QC ok. This shows overall precision, (analyser and extraction precision), was acceptable for the analysis of patient samples.
Note on F30: Hall Sebastian: Lvl 1 QC accepted as both sets of extracted QC are within limits showing acceptable overall precison (analyser and extraction precision). Cal'd before next run due to drift down of QC.
Note on F60: Hall Sebastian: Level 1 QC not rerun by immunology as they were still using the manufacturers QC ranges due to biochem not communicating the new ranges. This has now been communicated. QC accepted as only 1 level, still within 3SD and extracted QC ok. This shows overall precision, (analyser and extraction precision), was acceptable for the analysis of patient samples.

## LOT CT2BA and CT5BG
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 19.8571 | 186.5714
|  | Lot No. |  | Month | FEBRUARY-APRIL | establish current target and range using 5% target SD and complete |  |  | SD | 0.99286 | 9.3286
Level 1 |  | CT2BA |  | Year | 2024 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 17.8714 | 167.9143
Level 2 |  | CT5BG |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 21.8429 | 205.2286
Date | QC Level 1 | QC Level 2 | Comments
| 18-22 | 168-205 | [9 empty] | Level 1 | CT2BA |  |  |  | Level 2 | CT5BG
target | 20 | 187 | [9 empty] | target mean | -1SD | +1SD | -2SD | +2SD | target mean | -1SD | +1SD | -2SD | +2SD
2024-02-06 | 18 | 198 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-02-08 | 22 | 179 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-02-15 | 17 | 189 | Level 1 borderline, accepted by SH/CT. Calibrate next run |  |  | TREND REVIEW | [5 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-02-20 | 19 | 186 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-02-22 | 21 | 183 |  |  |  | 2024-02-01 | [5 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-02-27 | 22 | 192 |  |  |  | Range | QC 1 | QC 2 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-02-29 | 20 | 179 |  |  |  | Mean | 20 | 188.0909 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-03-07 | 21 | 172 |  |  |  | SD | 1.7321 | 9.0825 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-03-12 | 22 | 194 |  |  |  | % CV | 8.6603 | 4.8288 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-03-14 | 19 | 197 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-03-19 | 19 | 200 |  |  |  | 2024-03-01 | [5 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-03-21 | 22 | 194 |  |  |  | Range | QC 1 | QC 2 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-03-26 | 21 | 177 |  |  |  | Mean | 20.3889 | 188.6111 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-04-04 | 21 | 190 |  |  |  | SD | 1.577 | 7.942 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-04-10 | 20 | 186 |  |  |  | %CV | 7.7346 | 4.2108 |  |  |  | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-04-11 | 22 | 192 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-04-16 | 19 | 196 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
2024-04-18 | 22 | 191 | [9 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286
[12 empty] | 19.8571 | 18.8643 | 20.85 | 17.8714 | 22.0029 | 186.5714 | 177.2429 | 195.9 | 167.9143 | 205.2286

## Lot CT5BF and CT2B6
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Range | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 22.1818 | 180.5455 | Mean | 22.1818 | 180.5455
| Received | Lot No. | Expires | Month | DEC/JAN | establish current target and range using 5% target SD and complete |  |  | SD | 1.1091 | 9.0273 | SD | 1.1091 | 9.0273
Level 1 |  | CT2B6 |  | Year | 2023/2024 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 19.9636 | 162.4909 | lower limit | 19.9636 | 162.4909
Level 2 |  | CT5BF |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 24.4 | 182.7636 | upper limit | 24.4 | 182.7636
Date | QC Level 1 | QC Level 2 | Comments
| 0-40 | 109.8-256.2
target | 20 | 183
2023-12-22 | 24 | 180 | Accepted by RH, calibrate next run
2023-12-28 | 23 | 180
| 24 | 178 | Calibrated on 28th December no change. |  |  |  | TREND REVIEW
2024-01-09 | 22 | 185
2024-01-11 | 22 | 193
Date | QC Level 1 | QC Level 2 |  |  |  |  | Range | QC 1 | QC 2
| 20-25 | 166-202 |  |  |  |  | Mean | 23 | 183.2
target | 23 | 183 |  |  |  |  | SD | 1 | 6.0581
2024-01-15 | 24 | 197 |  |  |  |  | % CV | 4.3478 | 3.3068
2024-01-18 | 21 | 184
2024-01-23 | 24 | 193
2024-01-25 | 20 | 166 |  |  |  |  | Range | QC 1 | QC 2
2024-01-31 | 17 | 147 | Accepted by RH. To be cal'd next run |  |  |  | Mean | 21.2 | 177.4
[7 empty] | SD | 2.9496 | 20.7678
[7 empty] | %CV | 13.9131 | 11.7067
[7 empty] | Range | QC 1 | QC 2
[7 empty] | Mean | 22.1818 | 180.5455
[7 empty] | SD | 2.1826 | 14.0097
[7 empty] | %CV | 9.8395 | 7.7597
[7 empty] | Ranges | 24.4 | 198.6
[8 empty] | 19.9636 | 162.4909

## Lot CT5BE and CT2B6
FCAL Internal QC Log - Phadia | [5 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Range | QC 1 | QC 2
[6 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 19.1429 | 200.8333 | Mean | 19.2353 | 192.7059
| Received | Lot No. | Expires | Month | Oct | establish current target and range using 5% target SD and complete |  |  | SD | 0.95714 | 10.0417 | SD | 0.96176 | 9.6353
Level 1 |  | CT2B6 |  | Year | 2023 | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 17.2286 | 180.75 | lower limit | 17.3118 | 173.4353
Level 2 |  | CT5BE |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 21.0571 | 220.9167 | upper limit | 21.1588 | 211.9765
[9 empty] | Ranges applied 18/7/23 and given to imm |  |  | Ranges adjusted 7/9/23 and given to immunology
Date | QC Level 1 | QC Level 2 | Comments
| 17-21 | 175-214
target | 20 | 196
2023-10-03 | 20 | 217 | accepted by CT, to calibrate before next run
2023-10-05 | 18 | 191
2023-10-13 | 22 | 209 | accepted by SH |  |  |  | TREND REVIEW
2023-10-17 | 22 | 209 | accepted by RH
2023-10-24 | 20 | 205
2023-10-26 | 19 | 189 |  |  |  |  | Range | QC 1 | QC 2
2023-10-31 | 19 | 207 |  |  |  |  | Mean | 19.2727 | 197.8182
2023-11-02 | 18 | 189 |  |  |  |  | SD | 1.6181 | 12.0235
2023-11-06 | 18 | 180 |  |  |  |  | % CV | 8.3957 | 6.078
2023-11-09 | 17 | 185
2023-11-14 | 19 | 195
2023-11-16 | 22 | 190 | accepted by SH |  |  |  | Range | QC 1 | QC 2
2023-11-21 | 21 | 190 | Accepted by RH ? Calib next week. |  |  |  | Mean | 19.4 | 193.95
2023-11-23 | 17 | 197 |  |  |  |  | SD | 1.6351 | 12.0546
2023-11-28 | 21 | 212 |  |  |  |  | %CV | 8.4286 | 6.2153
2023-11-30 | 19 | 177
2023-12-05 | 20 | 192
2023-12-07 | 19 | 182
2023-12-12 | 17 | 177 |  |  |  |  | Range | QC 1 | QC 2
2023-12-14 | 20 | 186 |  |  |  |  | Mean | 19.4 | 193.95
[7 empty] | SD | 1.6351 | 12.0546
[7 empty] | %CV | 8.4286 | 6.2153

## Lot CT5BE and CT2B5
FCAL Internal QC Log - Phadia | [6 empty] | What to do when lot number of Immunology QC changes: |  |  | Range | QC 1 | QC 2 | Range | QC 1 | QC 2
[7 empty] | Manufactures range to be used as a guide until > 5 replicates available to |  |  | Mean | 19 | 197.3333 | Mean | 17.9444 | 194.5
| Received | Lot No. | Expires | Month | March |  | establish current target and range using 5% target SD and complete |  |  | SD | 0.95 | 9.8667 | SD | 0.89722 | 9.725
Level 1 |  | CT2B5 |  | Year | 2023 |  | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  | lower limit | 17.1 | 177.6 | lower limit | 16.15 | 175.05
Level 2 |  | CT5BE |  |  |  |  | to update their analyser Levy Jennings plot |  |  | upper limit | 20.9 | 217.0667 | upper limit | 19.7389 | 213.95
[10 empty] | Ranges applied 18/7/23 and given to imm |  |  | Ranges adjusted 7/9/23 and given to immunology
Date | QC Level 1 | QC Level 2 | Comments
| 16-20 | 175-214
target | 20 | 196
2023-06-13 | 19 | 197 | New LOT Pos QC
14.6.23 | 18 | 195 | New LOT Pos QC | Neg QC
22.06.23 | 14 | 164 | [5 empty] | TREND REVIEW
2023-06-27 | 13 | 147 | calibrate next run.
2023-07-06 | 18 | 197 | calibration |  |  |  |  | JUNE/JULY
2023-07-11 | 18 | 208 | [5 empty] | Range | QC 1 | QC 2
2023-07-13 | 19 | 191 | [5 empty] | Mean | 17.4545 | 188.5455
2023-07-18 | 21 | 232 | Calibrate next run (Immunology to use established ranges rather than manufacturers) |  |  |  |  | SD | 2.6216 | 22.4827
2023-07-20 | 20 | 184 | [5 empty] | % CV | 15.0195 | 11.9243
2023-07-25 | 14 | 173 | calibrate again on next run.
2023-07-27 | 18 | 186 | [5 empty] | JUNE/JULY/AUGUST
2023-08-01 | 17 | 185 | [5 empty] | Range | QC 1 | QC 2
2023-08-03 | 17 | 175 | [5 empty] | Mean | 17.5 | 190.8636
08//08/23 | 18 | 198 | [5 empty] | SD | 2.1101 | 17.6022
10.08.23 | 17 | 182 | [5 empty] | %CV | 12.0575 | 9.2224
2023-08-15 | 18 | 205
2023-08-17 | 15 | 190 | QC1 running low, QC accepted. SH |  |  |  |  | July/August/September
2023-08-22 | 18 | 205
2023-08-24 | 16 | 195 | [5 empty] | Range | QC 1 | QC 2
2023-08-29 | 17 | 190 | [5 empty] | Mean | 17.96 | 192.48
2023-08-31 | 19 | 186 | [5 empty] | SD | 1.6703 | 13.9527
2023-09-05 | 21 | 214 | QC accepted RH |  |  |  |  | %CV | 9.3003 | 7.2489
2023-09-07 | 19 | 180
2023-09-12 | 16 | 185
2023-09-14 | 18 | 179
2023-09-19 | 18 | 172 | re ran 173 accepted by RH |  |  | Calibrate next run
2023-09-21 | 20 | 200
2023-09-26 | 19 | 204
2023-09-28 | 18 | 196

## Lot CT5BC and CT2B3
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | March
Level 1 |  | CT2B3 |  | Year | 2023
Level 2 |  | CT5BC
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology QC changes: | Range | QC 1 | QC 2
| 18-23 | 162-198 | manufacturers ranges |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 20.5 | 180.2
target | 20 | 196 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1.025 | 9.01
2023-03-23 |  | 172 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 18.45 | 162.18
30.03.23 | 19 | 160 |  |  |  | to update their analyser Levy Jennings plot |  | 22.55 | 198.22
2023-04-04 | 21 | 183
2023-04-06 | 22 | 190 | Tell Immunology new ranges today!
2023-04-13 | 22 | 185
2023-04-18 | 24 | 216 | Run accepted as extracted IQC is within range for all runs. Immunology only received new ranges after this run was completed.
20.04.23 | 18 | 170
2023-04-25 | 23 | 206 | RERUN: 25/201 accepted by CT |  |  |  | may/june
2023-04-27 | 18 | 163 |  |  |  |  | Range | QC 1 | QC 2
02.05.23 | 21 | 192 |  |  |  |  | Mean | 21.3 | 191.6
04.05.23 | 19 | 178 |  |  |  |  | SD | 2.3118 | 15.7
09.05.23 | 22 | 194 |  |  |  |  | % CV | 10.8535 | 8.1941
2023-05-11 | 21 | 184
16.05.23 | 22 | 194
30.05.23 | 19 | 193
2023-06-01 | 18 | 173
06.06.23 | 23 | 197 | sl high both levels.
2023-06-08 | 22 | 181
2023-06-13 | 26 | 230 | Both levels of QC high. RR Neg QC 19 and RR New LOT Pos QC: 197 (see new tab), both acceptable. A selection of patient samples were also rerun and showed no significant difference.

## Lot CT2B2 and CT5BB
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2B2 |  | Year | 2022
Level 2 |  | CT5BB
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology QC changes: | Range | QC 1 | QC 2
| 16-20 | 170-208 | manufacturers ranges |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 18 | 189.2
target | 20 | 196 |  |  |  | establish current target and range using 5% target SD and complete | SD | 0.9 | 9.46
2022-12-13 | 19 | 187 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 16.2 | 170.28
2022-12-15 | 18 | 190 |  |  |  | to update their analyser Levy Jennings plot |  | 19.8 | 208.12
2022-12-20 | 17 | 196
2022-12-22 | 16 | 177
2022-12-29 | 17 | 178
2023-01-03 | 17 | 154 | re-ran patients on next run, no significant change overall, ok to report
5.1.23 | 16 | 210 | post-calibration. only slightly high, see comment above. Accepted by CT
10.01.23 | 17 | 198
2023-01-11 | 18 | 197
17.1.23 | 18 | 179
2023-01-18 | 18 | 182
24.01.23 | 18 | 180
26.01.23 | 18 | 188
2023-01-31 | 19 | 190
2023-02-01 | 19 | 191
2023-02-07 | 17 | 184
2023-02-09 | 17 | 173
2023-02-14 | 20 | 208
2023-02-16 | 17 | 186
2023-02-21 | 20 | 198
2023-02-23 | 17 | 187
2023-02-27 | 17 | 219 | slightly high accepted.
2023-03-02 | 16 | 172
2023-03-07 | 22 | 225
2023-03-07 | 19 | 211 | pp accepted qc Calibration to be performed.
2023-03-09 | 18 | 185 | post calibration
2023-03-14 | 20 | 171
2023-03-16 | 18 | 188
2023-03-23 | 17 | 167 | Positive is slightly low but accepted
2023-03-28 | 18 | 199

## Lot CT2B0 and CT5BA
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2B0 |  | Year
Level 2 |  | CT5BA
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology QC changes: | Range | QC 1 | QC 2
| 14-19 | 193 - 233 | manufacturers range for L1 |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 16.8333
2022-09-16 | 17 | 197 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1.169
2022-09-22 | 18 |  |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 14-19
2022-09-27 | 17 | 196 |  |  |  | to update their analyser Levy Jennings plot
2022-10-04 | 15 | 198
2022-10-06 | 16 | 198 | new range introduced for L1
2022-10-11 | 18 | 207 | manufacturers range for L1 = <40
2022-10-13 | 17 | 204
18.10.22 | 20 | 229
2022-10-20 | 22 | 221
2022-10-25 | 19 | 230
1900-01-26 02:24 | 18 | 209
2022-11-01 | 17 | 201
2022-11-03 | 15 | 206
8.11.22 | 18 | 209
2022-11-10 | 19 | 213
15.11.22 | 19 | 220
2022-11-17 | 20 | 232
22.11.22 | 18 | 194
2022-11-24 | 19 | 201
2022-11-29 | 17 | 205
2022-12-01 | 19 | 193
2022-12-06 | 17 | 205
2022-12-08 | 18 | 211

## LOT CT2AZ and CT5BA
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AZ |  | Year
Level 2 |  | CT5BA
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology QC changes: | Range | QC 1 | QC 2
| 19-23 | 111-260 | manufacturers range for L2 |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 20.25 | 213
[6 empty] | establish current target and range using 5% target SD and complete | SD | 1.0125 | 10.65
2022-08-30 | 18 | 192 | within manufacturers limits |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 19-23 | 193-233
2022-09-01 | 21 | 204 |  |  |  | to update their analyser Levy Jennings plot
2022-09-05 | 19 | 218
2022-09-08 | 21 | 234
2022-09-13 | 23 | 227
15.09.22 | 18 | 203 | check if still using manufacturrer range
NEW RANGE ASSIGNED LEVEL 2 193-233
Mean | 20 | 213
SD | 2 | 16.025
%CV | 10 | 7.5235

## LOT CT2AZ and CT5B8
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AZ |  | Year
Level 2 |  | CT5B8
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology QC changes: | Range | QC 1 | QC 2
| 19-23 | 192-235 | manufacturers range for L1 |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 20.5 | 213.4
|  |  |  | <40 |  | establish current target and range using 5% target SD and complete | SD | 1.025 | 10.67
[6 empty] | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 19-23 | 192-235
to update their analyser Levy Jennings plot
12.05.22 | 20 | 224
2022-05-17 | 20 | 222
2022-05-24 | 20 | 207
26.05.22 | 19 | 196
31.5.22 | 23 | 204
2022-06-07 | 21 | 201
09.06.22 | 23 | 201
14.6.22 | 22 | 223
2022-06-16 | 20 | 217
21.6.22 | 22 | 222
23.6.22 | 19 | 206
2022-06-28 | 20 | 207
29/06/22 | 21 | 228
2022-07-05 | 20 | 214
2022-07-07 | 18 | 193
2022-07-12 | 19 | 212
2022-07-19 | 23 | 238 | L2 sl high but not reran, talk to immunology
2022-07-26 | 23 | 206
2022-07-28 | 21 | 219
2022-08-02 | 21 | 208
2022-08-04 | 20 | 229
2022-08-09 | 19 | 230
18/08/202 | 19 | 211
2022-08-23 | 19 | 199
2022-08-25 | 18 |  | acceptable for immunology
Mean | 20.4 | 213.2083
SD | 1.5546 | 11.9491
%CV | 7.6204 | 5.6044

## LOT CT2AY and CT5B7
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AY |  | Year
Level 2 |  | CT5B7
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 20-24 | 111-259 | manufacturers range for L2 |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean
2022-01-18 | 21 | 199 |  |  |  | establish current target and range using 5% target SD and complete | SD
2022-01-27 | 21 | 224 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range
2022-01-28 | 21 | 210 |  |  |  | to update their analyser Levy Jennings plot
2022-02-01 | 22 | 200
2022-02-03 | 22 | 216
2022-02-08 | 20 | 211
2022-02-15 | 20 | 231
21.2.22 | 21 | 225
24.2.22 | 19 | 209
2022-03-01 | 22 | 234
2022-03-03 | 21 | 215
2022-03-08 | 21 | 202
2022-03-10 | 21 | 223
15.03.22 | 21 | 204
22.03.22 | 21 | 208
2022-03-29 | 20 | 233
31.03.22 | 20 | 195
05.04.22 | 18 | 183 | re run L1 | L1 QC still low but within manufacturers range
| 17-21 | 111-259 |  | Range adjusted 21/12/21 previous range 17-21
2022-04-12 | 19 | 195
14.04.22 | 18 | 193
21.04.22 | 24 | 226
26.04.22 | 19 | 200
28.04.22 | 21 | 224
2022-05-03 | 21 | 211
Mean | 20.5833 | 211.2917
SD | 1.3486 | 13.9953
%CV | 6.5521 | 6.6237

## LOT CT2AY and CT5B8
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AY |  | Year
Level 2 |  | CT5B8
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 20-24 | 192-235 | manufacturers range for L2 |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean
[6 empty] | establish current target and range using 5% target SD and complete | SD
2022-05-05 | 23 | 229 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range
2022-05-10 | 22 | 205 |  |  |  | to update their analyser Levy Jennings plot
12.05.22 | 20 | 224
2022-05-17 | 20 | 222
Mean | 21.25 | 220
SD | 1.5 | 10.4243
%CV | 7.0588 | 4.7383

## LOT CT2AY and CT5B6
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AY |  | Year
Level 2 |  | CT5B6
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 20-24 | 110-256 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 19 | 203
19.10.21 | 19 | 197 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1 | 10
21.10.21 | 19 | 205 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 17 - 21 | 183 - 223
26.10.21 | 21 | 201 |  |  |  | to update their analyser Levy Jennings plot |  | 17
28.10.21 | 19 | 210 | Target assigned to L2 |  |  |  |  | 21
2021-11-02 | 22 | 208
4.11.21 | 21 | 211
2021-11-09 | 21 | 197
15.11.21 | 23 | 216 | ? Lvl 1 QC shift
2021-11-18 | 18 | 196
2021-11-23 | 18 | 184
2021-11-25 | 18 | 173 | Checked ranges with immunology
2021-11-30 | 17 | 203
01.12.21 | 21 | 136 |  |  |  | j
2021-12-07 | 24 | 235
2021-12-09 | 21 | 195
14.12.21 | 23 | 220
2021-12-16 | 24 | 216
2021-12-21 | 24 | 228 |  |  | lvl 1 new range 22±2
2021-12-30 | 23 | 206
04.01.22 | 24 | 208
11.01.22 | 20 | 208
13.01.22 | 22 | 238 | Checked ranges with immunology - Phadia range = 110-256 RH will sort
18.01.22 | 21 | 199
2022-01-27 | 21 | 224
2022-02-03 | 22 | 216
2022-02-08 | 20 | 211
2022-02-10 | 20
Mean | 20.963 | 205.4231
SD | 2.0472 | 20.1776
%CV | 9.7656 | 9.8224
.

## LOT CT2AY and CT5B4
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AY |  | Year
Level 2 |  | CT5B4
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 17-21 | 163-203 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 18.7143
2021-09-21 | 17 |  |  |  |  | establish current target and range using 5% target SD and complete | SD | 0.93571
21.9.21 | 20 | 176 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 17-21
23.09.21 | 15 | 189 |  |  |  | to update their analyser Levy Jennings plot |  | 16.8429
2021-09-28 | 20 | 185 | [5 empty] | 20.5857
2021-09-30 | 19 | 194
2021-10-05 | 20 | 198 | Rpt 168
2021-10-07 | 20 | 196 | Assigned target to L1
2021-10-12 | 19 | 195
2021-10-15 | 19 | 193
19.10.21 | 19 | 197 | new lot CT5B6
Mean | 18.8 | 191.4444
SD | 1.6193 | 7.0907
%CV | 8.6134 | 3.7038
.

## LOT CT2AW and CT5B4
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AW |  | Year
Level 2 |  | CT5B4
Date | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 22-26 | 163-203 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 23.375 | 183.0769
2021-08-17 | 22 | 159 | Will calibrate before next batch |  |  | establish current target and range using 5% target SD and complete | SD | 1.6863 | 20.2328
2021-08-17 | 22 | 162 | ranges for both remain same as when this lot previously used |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 22-26 | 163-203
19.8.21 | 23 | 190 |  |  |  | to update their analyser Levy Jennings plot
24.08.21 | 25 | 195
2021-09-02 | 23 | 191
7.9.21 | 23 | 194
9.9.21 | 24 | 202 | Means adjusted based on data from june/july
16.9.21 | 22 | 201
Mean | 23 | 186.75
SD | 1.069 | 16.7652
%CV | 4.648 | 8.9773
.

## LOT CT2AX and CT5B4
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AX |  | Year
Level 2 |  | CT5B4
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 22-26 | 163-199 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 23.6 | 186
2021-06-29 | 23 | 187 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1.18 | 20.379
2021-07-01 | 24 | 182 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 22-26 | 163-199
2021-07-06 | 25 | 190 |  |  |  | to update their analyser Levy Jennings plot
2021-07-09 | 26 | 193
2021-07-13 | 23 | 176
2021-07-15 | 25 | 196
20.07.21 | 22 | 170
2021-07-21 | 25 | 194
2021-07-27 | 25 | 186
2021-08-03 | 22 | 162 | check ranges ~(QC2) with immunology
2021-08-05 | 26 | 196
10.8.21 | 23 | 177
2021-08-12 | 22 | 171
Mean | 23.9231 | 183.0769
SD | 1.4979 | 11.079
%CV | 6.2612 | 6.0515
.

## LOT CT2AW and CT5B2
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AW |  | Year
Level 2 |  | CT5B2
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 22-26 | 163-199 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 23.6 | 182.6
19.01.21 |  | 183 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1.18 | 9.13
2021-01-21 |  | 180 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 22-26 | 163-199
2021-01-26 |  | 194 |  |  |  | to update their analyser Levy Jennings plot
2021-01-28 |  | 167
02.02.21 |  | 184
04.02.21 |  | 188
09.02.21 |  | 175
11.02.21 |  | 173
16.02.21 |  | 158
18.02.21 |  | 206
23.02.21 |  | 181
25.02.21 |  | 161
2021-03-02 |  | 167
09.03.21 |  | 161 | (158, rpt) |  | QC range adjusted from 166--202 (163-199)
11.03.21 |  | 191
2021-03-16 |  | 190
2021-03-18 |  | 186
23.3.21 |  | 191
25.3.21 |  | 178
11.5.21 |  | 193 |  | LOT restarted in May
13.05.21 |  | 163
2021-05-18 |  | 175
2021-05-20 |  | 186
2021-05-25 | 24 | 178 | LOT CT2AW running in
2021-05-27 | 23 | 164
2021-06-01 | 24 | 177
2021-06-03 | 24 | 183
2021-06-11 | 23 | 178 | QC range assigned 22-26
2021-06-16 | 23 | 177
2021-06-22 | 23 | 169
2021-06-24 | 23 | 169
Mean | 23.375 | 178.2778
SD | 0.51755 | 11.4309
%CV | 2.2141 | 6.4118
.

## LOT CT2AV and CT5B4
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | March
Level 1 |  | CT2AV |  | Year | 2021
Level 2 |  | CT5B4 | 2022-01-31
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 19-23 | 170-207 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 22 | 189.2
25.3.21 | 20 | 178 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1.1 | 9.46
30.3.21 | 23 | 206 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 19-23 | 170-207
2021-04-01 | 25 | 190 |  |  |  | to update their analyser Levy Jennings plot |  |  | 207.92
2021-04-08 | 20 | 181 | [6 empty] | 170.28
12.04.21 | 22 | 183
16.04.21 | 21 | 186
2021-04-20 | 22 | 171
22.4.21 | 20 | 186
27.4.21 | 21 | 171
4.5.21 | 24 | 214
2021-05-05 | 20 | 180
Mean | #DIV/0! | #DIV/0!
SD | 1.7477 | 13.3566
%CV | #DIV/0! | #DIV/0!
.

## LOT CT2AV and CT5B2
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AV |  | Year
Level 2 |  | CT5B2
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: | Range | QC 1 | QC 2
| 19-23 | 158-193 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to | Mean | 21 | 182.6
19.01.21 | 20 | 183 |  |  |  | establish current target and range using 5% target SD and complete | SD | 1.05 | 9.13
2021-01-21 | 23 | 180 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them | range | 19-23 | 163-199
2021-01-26 | 20 | 194 |  |  |  | to update their analyser Levy Jennings plot
2021-01-28 | 20 | 167
02.02.21 | 22 | 184
04.02.21 | 20 | 188
09.02.21 | 20 | 175
11.02.21 | 17 | 173
16.02.21 | 18 | 158
18.02.21 | 24 | 206
23.02.21 | 22 | 181
25.02.21 | 19 | 161
2021-03-02 | 20 | 167
09.03.21 | 19 | 161 | (158, rpt) |  | QC range adjusted from 166--202 (163-199)
11.03.21 | 21 | 191
2021-03-16 | 21 | 190
2021-03-18 | 21 | 186
23.3.21 | 21 | 191
25.3.21 | 20 | 178
11.5.21 | 21 | 193 |  | LOT restarted in May
13.05.21 | 19 | 163
2021-05-18 | 19 | 175
2021-05-20 | 20 | 186
2021-05-25 | 24 | 178 | LOT CT2AW
Mean | 20.5455 | 181.0909
SD | 1.6934 | 12.3005
%CV | 8.2424 | 6.7925
.

## LOT CT2AU and CT5B2
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AU |  | Year
Level 2 |  | CT5B2
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes: |  |  |  | Range | QC 1 | QC 2
| 19-24 | 166-202 |  |  |  | Manufactures range to be used as a guide until > 5 replicates available to |  |  |  | Mean | 22 | 184.4
2020-10-20 | 21 | 198 |  | running in |  | establish current target and range using 5% target SD and complete |  |  |  | SD | 1.1 | 9.22
2020-10-27 | 24 | 190 | RR LEVEL 1 = 20 |  | running in | form in folder. Running mean +/- 2*5%. Tell immunology for them |  |  |  | range | 19-24 | 166-202
2020-10-28 | 19 | 154 |  |  | running in | to update their analyser Levy Jennings plot
2020-11-19 | 23 | 190 |  |  | running in
2020-11-23 | 23 | 190 | Accepted by EH |  | running in
2020-12-01 | 23 | 186 | lot specific ranges assigned at 5% CV
2020-12-08 | 21 | 167
2020-12-10 | 23 | 185
2020-12-15 | 21 | 191
2020-12-17 | 24 | 171
2020-12-22 | 24 | 194
2020-12-24 | 22 | 172
2020-12-31 | 21 | 153
2021-01-05 |  | 164
2021-01-12 | 22 | 173 |  |  |  |  | Range | QC 1 | QC 2
2021-01-14 | 22 | 178 |  |  |  |  | Mean | 22 | 172
2021-01-19 | 20 | 183 |  |  |  |  | SD | 1.1 | 8.6
2021-01-21 | 23 | 180 |  |  |  |  | range | 19-24 | 163-181
2021-01-26 |  | 194
2021-01-28 | 20 | 167
2021-02-02 | 22 | 184
2021-02-05 | 20 | 188
Mean | 21.8182 | 177
SD | 1.4709 | 11.9583
%CV | 6.7418 | 6.7561
.

## LOT CT2AT and CT5AZ
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AT |  | Year
Level 2 |  | CT5AZ
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes:
| 18-23 | 113-263 | Manufacturer's range for CT2AT |  |  | Manufactures range to be used as a guide until > 5 replicates available to
04.08.20 | 18 | 169 |  |  |  | establish current target and range using 5% target SD and complete
11.08.20 | 18 | 170 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them
12.08.20 | 19 | 187 |  |  |  | to update their analyser Levy Jennings plot
20.08.20 | 21 | 206
25.08.20 | 20 | 199
2020-09-01 | 20 | 207
2020-09-08 | 21 | 212
2020-09-10 | 21 | 205
2020-09-14 | 21 | 176
2020-09-22 | 23 | 178
Mean | 23 | 178
SD | #DIV/0! | #DIV/0!
%CV | #DIV/0! | #DIV/0!
.

## Current CHART
FCAL QC Phadia
Material
| Received | Lot No. | Expires | In Use | Month | January - December
Level 1 |  | CT2AP |  | 2019-06-13 | Year | 2019
Level 2 |  | CT5AX |  | 2019-06-13

## LOT CT2AS AND CT5AZ
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AS |  | Year | 2019
Level 2 |  | CT5AZ
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes:
| 16-23 | 113-263 | Manufacturer's range for CT2AS = 16-23 |  |  | Manufactures range to be used as a guide until > 5 replicates available to
|  | 168 | QC low |  |  | establish current target and range using 5% target SD and complete
|  | 179 | repeated QC |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them
|  | 174 |  |  |  | to update their analyser Levy Jennings plot
171
31.12.19 | 20 | 172
2020-01-07 | 22 | 198
2020-01-13 | 17 | 171
2020-01-16 | 18 | 165 | l1 RPT, 18
2020-01-24 | 18 | 171 | L1 RPT,18
2020-01-28 | 18 | 179 | L1 RPT 19
2020-01-30 | 19 | 174
2020-01-31 | 18 | 165
2020-02-06 | 19 | 178
2020-02-11 | 22 | 209
2020-02-13 | 19 | 177
18.02.20 | 21 | 192
2020-02-21 | 19 | 177
2020-02-25 | 20 | 199
26/20/2020 | 20 | 184
03.03.20 | 21 | 204
2020-03-04 | 20 | 182
2020-03-05 | 20 | 182
2020-03-09 | 19 | 176
2020-03-12 | 20 | 181 | ok
2020-03-17 | 22 | 197 | ok
2020-03-19 | 21 | 185
2020-03-24 | 21 | 184
2020-04-02 | 19 | 177
2020-04-21 | 21 | 208
2020-04-30 | 21 | 201
2020-05-07 | 19 | 200
2020-05-12 | 17 | 157
19.05.20 | 18 | 176
26.05.20 | 19 | 180
02.06.20 | 19 | 190
06.06.20 | 15 | 156
11.06.20 | 16 | 154
30.06.20 | 20 | 186
02.07.20 | 20 | 188
07.07.20 | 20 | 171
14.07.20 | 22 | 192
16.07.20 | 21 | 193
17.07.20 | 19 | 181
21.07.20 | 21 | 186
23.07.20 | 20 | 186
Mean | 19.6111 | 183.5278
SD | 1.5726 | 13.3384
%CV | 8.0188 | 7.2678
.

## LOT CT2AR AND CT5AZ
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AR |  | Year | 2019
Level 2 |  | CT5AZ
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes:
| 20-27 | 113-263 | Manufacturer's range for CT5AZ = 113-263 |  |  | Manufactures range to be used as a guide until > 5 replicates available to
2019-12-17 | 19 | 168 | QC low |  |  | establish current target and range using 5% target SD and complete
2019-12-17 | 20 | 179 | repeated QC |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them
2019-12-19 | 20 | 174 |  |  |  | to update their analyser Levy Jennings plot
2019-12-24 | 21 | 171
Mean | 20 | 173
SD | 0.8165 | 4.6904
%CV | 4.0825 | 2.7112

## LOT CT2AR AND CT5AY
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AR |  | Year | 2019
Level 2 |  | CT5AY
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes:
| 20-27 | 195 - 238 | Manufacturer's range for CT2AR = 20-27 |  |  | Manufactures range to be used as a guide until > 5 replicates available to
2019-10-31 |  | 206 |  |  |  | establish current target and range using 5% target SD and complete
2019-11-05 |  | 223 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them
2019-11-07 |  | 216 |  |  |  | to update their analyser Levy Jennings plot
2019-11-12 |  | 226
2019-11-14 |  | 213
2019-11-19 | 23 | 207
2019-11-21 | 24 | 215
2019-11-26 | 25 | 226
2019-11-28 | 22 | 213
2019-12-03 | 20 | 187
2019-12-03 |  | 208 | repeated
2019-12-10 | 21 | 172
2019-12-12 | 20 | 197
2019-12-12 | 21 |  | repeated
2019-12-17 | 19 | 168 | low qc |  |  | Lot number change QC 2 CT5AZ
2019-12-17 | 20 | 179 | repeated
Mean | 21.5 | 203.7333
SD | 1.9579 | 19.0243
%CV | 9.1065 | 9.3378

## LOT CT2AN AND CT5AY
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month
Level 1 |  | CT2AN |  | Year | 2019
Level 2 |  | CT5AY
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes:
| 20-27 | 195 - 238 | Manufacturers range: 122.9-286.7 (mean 205) |  |  | Manufactures range to be used as a guide until > 5 replicates available to
2019-10-31 | 23 | 206 | New Lot Level 2 QC , Level 1 same lot |  |  | establish current target and range using 5% target SD and complete
2019-11-05 | 25 | 223 | Need to review Level 2 range after run for few days |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them
2019-11-07 | 24 | 216 |  |  |  | to update their analyser Levy Jennings plot
2019-11-12 | 24 | 226
2019-11-14 | 24 | 213
2019-11-19 |  | 207 | New target ranges established see report in FCAL folder
Mean | 24 | 215.1667
SD | 0.70711 | 8.1833
%CV | 2.9463 | 3.8032

## LOT CT2AN AND CT5AW
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | January - December
Level 1 |  | CT2AN |  | Year | 2019
Level 2 |  | CT5AW
Date/Range | QC Level 1 | QC Level 2 | Comments
| 20-27 | 202-223
2019-09-10 | 22 | 205
2019-09-12 | 20 | 210
2019-09-17 | 21 | 217
2019-09-19 | 21 | 221
2019-09-24 | 21 | 207
2019-09-26 | 21 | 200
2019-10-01 | 20 | 203
2019-10-10 | 25 | 249
2019-10-15 | 21 | 206
2019-10-22 | 21 | 214
2019-10-24 | 23 | 195 | Full re-calibration done
2019-10-29 | 22 | 202
2019-10-29 | 23 | 206
2019-10-31 | 23 | 206 | New Lot Level 2 QC
2019-11-05 | 25 | 223
Mean | 21.9333 | 210.9333
SD | 1.5796 | 13.0464
%CV | 7.202 | 6.1851

## LOT CT2AP AND CT5AW
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | January - December
Level 1 |  | CT2AP |  | Year | 2019
Level 2 |  | CT5AW
Date/Range | QC Level 1 | QC Level 2 | Comments
| 20-27 | 202-223
2019-08-29 | 25 | 225 | Borderline, accepted CT. Review next time
2019-09-03 | 23 | 220 | Full calibration done
2019-09-05 | 22 | 225 | Full calibration done
2019-09-10 | 22 | 205
2019-09-12 | 20 | 210
2019-09-17 | 21 | 217
2019-09-19 | 21 | 221
2019-09-24 | 21 | 207
2019-09-26 | 21 | 200
2019-10-01 | 20 | 203
2019-10-10 | 25 | 249
2019-10-15 | 21 | 206
2019-10-22 | 21 | 214
2019-10-24 | 23 | 195 | Full re-calibration done
2019-10-29 | 22 | 202
2019-10-29 | 23 | 206
2019-10-31 | 23 | 206 | New Lot Level 2 QC
2019-11-05 | 25 | 223
Mean | 22.1667 | 213
SD | 1.6179 | 12.7971
%CV | 7.2989 | 6.008

## LOT CT2AP AND CT5AX
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | January - December
Level 1 |  | CT2AP |  | Year | 2019
Level 2 |  | CT5AX
Date/Range | QC Level 1 | QC Level 2 | Comments
| 20-27 | 202-223
2019-06-13 | 21 | 177
2019-06-21 | 20 | 183
2019-06-25 | 22 | 192
2019-06-27 | 20 | 192
2019-07-02 | 23 | 202
2019-07-04 | 21 | 202
2019-07-09 | 22 | 194
2019-07-11 | 22 | 212
2019-07-16 | 21 | 204
2019-07-31 | 23 | 216
2019-08-01 | 22 | 211
2019-08-06 | 25 | 221 | Changed range - confirmed with immunology
2019-08-09 | 20 | 216
2019-08-13 | 23 | 214
2019-08-15 | 22 | 209
2019-08-20 | 22 | 214
2019-08-22 | 23 | 213
Mean | 21.9375 | 205.9375
SD | 1.3401 | 10.8718
%CV | 6.1087 | 5.2792

## LOT CT2AM and CT5AU
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | January - December
Level 1 |  | CT2AM |  | Year | 2018
Level 2 |  | CT5AU
Date/Range | QC Level 1 | QC Level 2 | Comments
| <40 | 115-268
1.11.18 | 20 | 182
5.11.18 | 22 | 178
2018-11-08 | 19 | 185
2018-11-13 | 19 | 178
2018-11-15 | 19 | 215
2018-11-20 | 19 | 175
2018-11-22 | 19 | 182
2018-11-27 | 19 | 172
2018-12-03 | 23 | 192
2018-12-04 | 20 | 180
2018-12-05 | 19 | 177
2018-12-11 | 18 | 172
2018-12-18 | 18 | 181
2019-01-03 | 19 | 187
2019-01-08 | 19 | 174
2019-01-11 | 20 | 188
2019-01-15 | 23 | 217
2019-01-17 | 21 | 220
2019-01-22 | 21 | 194
2019-01-24 | 23 | 211
2019-01-28 | 23 | 219
2019-01-31 | 23 | 227
2019-02-02 | 24 | 227
2019-02-07 | 24 | 234
2019-02-12 | 20 | 198
2019-02-14 | 23 | 217
2019-02-19 | 35 | 208
2019-02-21 | 24 | 215
2019-02-26 | 23 | 218 | NEW LOT CT5AW
2019-03-07 | 25 | 215 | 19-23 and 203-231 new target ranges
Mean | 21.5172 | 198.4828
SD | 3.366 | 19.9672
%CV | 15.6433 | 10.0599

## CT2AM and CT5AU CHART
FCAL QC Phadia
Material
| Received | Lot No. | Expires | In Use | Month | January - December
Level 1 | 0 | CT2AM | #REF! | 0 | Year | 2018
Level 2 | 0 | CT5AU | #REF! | 0

## LOT CT2AT and CT5B0.
FCAL Internal QC Log - Phadia
| Received | Lot No. | Expires | Month | November
Level 1 |  | CT2AT |  | Year
Level 2 |  | CT5B0
Date/Range | QC Level 1 | QC Level 2 | Comments |  |  | What to do when lot number of Immunology L2 QC changes:
|  | 113-263 | Manufacturer's range for CT2AT |  |  | Manufactures range to be used as a guide until > 5 replicates available to
2020-10-07 | 20 | 175 |  |  |  | establish current target and range using 5% target SD and complete
2020-10-15 | 19 | 162 |  |  |  | form in folder. Running mean +/- 2*5%. Tell immunology for them
to update their analyser Levy Jennings plot
Mean | #DIV/0! | #DIV/0!
SD | #DIV/0! | #DIV/0!
%CV | #DIV/0! | #DIV/0!
.